Research report – Microsal

Nițu Adrian - *Ing,* Universitatea Politehnica București

*Abstract*— Microsal represents the collaboration of multiple faculties with the purpose to create a micro-embedded device meant to alleviate symptoms of xerostomia and also collect patient data regarding the condition in an unobstrusive way.

*Index Terms*—Xerostomia, Android, Embedded, Bluetooth Low Energy,

# INTRODUCTION

S

alivation is a very important process that helps the body in digestion, hygiene, as well as taste and comfort. Compounds in the saliva begin the digestion process by starting to break down starch and fat. Water soluble chemicals, like sugars and salts, are spread through the mouth until they reach the taste buds. Saliva helps the act of chewing by lubricating the tongue and teeth and forms the bolus (the mass of food chewed and mixed with saliva), also having a key role in the act of swallowing as it lubricates the throat. Continuous salivation improves hygiene by maintaining a relatively stable level of acidity and reducing the number of bacteria. High levels of acidity are correlated with rapid tooth decay, cavities and halitosis. Salivary function varies based on physical aspects, such as time of day, and the act of chewing, and based on psychological aspects, such as stress, or thinking about food.

Xerostomia is a condition characterized by dryness of the mouth estimated to affect between one-fifth and one-third of the adult population, and can be triggered by numerous reasons: old age, medication (e.g. beta-blockers are known to cause a decrease of the salivary function) radiation exposure (patients with multiple neck and throat x-rays have been known to have the salivary functions permanently disrupted), Sjögren’s syndrome, even cerebrovascular accidents.

Electro-stimulation of the nerves is a technique that can be used to trigger muscle contraction and reflexes. However, special care must be taken to trigger only the nerves that increase salivary production and not to bring any discomfort (pain, muscle twitches, etc.) while requiring a minimum amount of energy to be expended. Previous studies provided a window into previous signals used to stimulate nerves, so we could derive some insight regarding voltage and current limitations, stimulation duration, and other signal parameters

# Related work

The paper *Neuroelectrostimulation in Treatment of Hyposalivation and Xerostomia in Sjögren’s Syndrome: A Salivary Pacemaker* introduces and follows the development of previous devices that treat xerostomia using neuro-electo-stimulation. The paper also introduces term definition and condition status.

The inflammation and infection issues that usually accompany xerostomia can be successfully managed using modern medicine techniques, however, a solution is needed that is as non-invasive as possible, and that increases the comfort of the patients.

It is noted that xerostomia affects mostly old people, almost never children. This fact might be related to neurological effects, perhaps reflecting the anticholinergic action of many drugs; An anticholinergic is a neuro-transmitter blocker. Also, interestingly, women are more commonly affected than men

This supports the theory that, in most cases, the salivary glands appear to be functional, and that merely the neurological system is disrupted. A stimulation of the neurological signal should be enough to resolve the issues for most people.

Therefore the neurological control of salivary secretion represents the basis for a novel therapeutic approach. The paper distinguishes 3 generations; Microsal is an advanced 3rd generation prototype.

## First generation

The first generation is represented by a single project: a clunky system with a tongue suppressor probe. Although large, expensive, and somewhat not portable (While the device could run on batteries, it’s size and the size of the probe made multi-tasking hard), the device proved the effectiveness of the neuro-elecro-stimulation procedure in alleviating symptoms. As the device gave promising results in proof-of principle clinical studies and did cause any local or systemic adverse effects, it was approved by the US Food and Drug Administration in 1988 (PMA No. P860067)

|  |
| --- |
| First-generation neuroelectrostimulation device consisted of a hand-held probe, tipped with stainless steel electrodes, and a console that housed a battery and the electronic signal-generating power source, the size and shape of which were similar to a video or CD player |

## 2nd generation

Building on the success of the first generation, but with the goal of reducing costs, and increasing the portability of the device, the second generation came to be. The most important aspect of the second generation was the reduced device size, which allowed everything to be mounted on a simple mouth-guard.

Second-generation removable device consists of 3 components: a miniaturized electronic stimulator that has a signal generator, power source, and conducting circuitry; an intra-oral removable appliance; an infrared remote control.

The second most important aspect that the second generation improved upon is that now, the device can be worn non-invasively and controlled remotely through the day.

|  |
| --- |
| The miniaturized electronic stimulator is mounted in a removable intraoral appliance (A), which is under remote control that activates the stimulator (B). This device is applied into the mouth in a noninvasive manner (C) |

## 3rd generation

The third generation takes the process of reducing size to an extreme, creating a device that is small enough to completely fit inside a tooth. This allows for for an implant-supported neuro-electrostimulating device that can be permanently applied into the oral cavity, screwed onto an osteo-integrated dental implant inserted in the third molar area.

The current research also suggests that electro-neuro-stimulation can have longer lasting effects, as in some patients, the salivary flow increased measurably even when the device was deactivated. This brings hope that after several years of treatment it is possible that normal salivary response could be restored completely.

|  |
| --- |
| Figure shows the implantation procedure and application of the device. Transmucosal exposure of mandibular bone (A) is followed by preparation of the implant bed in mandibular bone (B) and insertion of the dental root implant (C). The neuro-electrostimulating device is shown in its applicator (D) and mounted onto the root implant (E). A radiograph of the implant-supported device is shown (F). |

# Workshop & Diploma project presentation

## Hardware

The current iteration of this project has already been investigated in a diploma project and already presented at an embedded devices conference[2] and at a medical workshop [1] with plans to demonstrate crrent research stages.

This chapter will describe the current iteration and some of the design choices.

Microsal represents the research efforts of multiple faculties to create a third-generation implant-supported neuro-electrostimulation embedded device. This research report will only report on the hardware/firmware/software details regarding the project. As described in the related work section, the device’s main goal is to electrically stimulate the bucal and lingual nerve in order to produce the salivary response and increase the comfort of the patient. Additionally, besides the stimulation logic, the design also includes 2 passive sensors that are used to objectively evaluate the current conditions of the patient.

The information taken from the sensors will be used to determine if the patient’s condition is deteriorating or improving over time, providing valuable insight on the effectiveness of the treatment. This represents the next logical step for the evolution of the generational prototypes. Also, to facilitate communication to the device we have decided to use a chip that has a Bluetooth Low Energy component integrated inside. The NRF51822 chip, by Nordic Semiconductors fits the bill, being a highly advanced ARM microcontroller, with BLE functionality and extremely low power requirements

The high level design is described in the upper image

|  |
| --- |
| design  High level design of the Microsal Embedded Device. There are two main sensors, one for PH, the other for humidity connected to a high end micro-controller, which is also connected to the electrodes logic. |

## Waveform

In order to have the required results, with a minimum of energy expenditure we set out to find a signal pattern that produces a measurable response.

We started by describing all possible signal parameters (see the waveform explanatory image), and implementing the control hardware and software in order to enable all of them. We considered that a sinusoid approximation is enough to stimulate the neurons without harming nearby tissue. The assumption was supported by the experiment description at [3].

Experimentally (table 1) we determined that the patient responds favorably to a wide range of stimulation patterns, the one notable difference being when we used a combination of low pulse duration, low frequency, and low voltage. These results will be applied in the power optimization phase of the next iteration).

|  |
| --- |
| https://lh5.googleusercontent.com/1jv1Mn4wrW_ul0z6sF4hdNV1-L6twJXPiSRypZ04s6OYr-b3SIJo4OulNZ24cmrYmJROQaP-fNTGukbbyou9GjS4AMPG396uz49Eh_OcUzFechwGpwqFGU11WdzHG-rBta2Wu3P0mLT1a8hHjQ  Android User Interface for the application that will communicate with the NRF51822 chip. The application is responsible for the waveform generation pattern and collecting data. The application is not bound to a particular device, it can connect to any device that match the specification |

|  |
| --- |
| BoxSchema (1)  High level design of the prototype that was used in experimentally determining which values of the signal parameters are crucial |
| https://lh5.googleusercontent.com/R_mlPxQE5rKOZbcm3NNlfP7gbFwmOfb055dFlHPOZjYVCpeLbZaKANPbF_wF45eRAfKwkka3Y0M8V5mD8qnXozyCJqQdYNQQ9Y2Gxgk2qxw3Kmq-oRPtbTusIzwNI_wQ0VGoSSMLVhQMlps0Lg  Waveform pattern , composed of multiple sets of pulses. The sets of pulses are represented as blue and each individual pulse as green. The system works as well just in the presence of the positive front of each pulse. |
| 1. Experimental data  |  | **Pulse length** | **Frequency** | **Voltage** | **Voltage swing** | **Effect** | | --- | --- | --- | --- | --- | --- | | 1 | 4 ms | 0.1 Hz | 3.25 V | positive | N | | 2 | 4 ms | 1 Hz | 3 V | positive | P | | 3 | 10 ms | 1 Hz | 3 V | positive | P | | 4 | 10 ms | 2 Hz | 3 V | positive | P | | 5 | 4 ms | 4 Hz | 3 V | positive | P | | 6 | 4 ms | 8 Hz | 3 V | positive | P | | 7 | 10 ms | 8 Hz | 3.25 V | positive | P | | 8 | 10 ms | 8 Hz | 5 V | positive | P | | 9 | 10 ms | 8 Hz | 5 V | positive & negative | P | |
| Dimensions & casing The device, once soldered will be stored in a bio-compatible metal cylinder casing, designed to provide protection and support. The casing will be attached similarly of a dental implant, inside the bone of the patient (3rd generation neuroelectrostimulation device). The casing is designed to provide the support and function of a real tooth, while not obstructing wireless communication. |
| C:\Users\Dan\Dropbox\eagle\Sparrow_BLE\eUp\top-bottom.png  PCB design of the Microsal Embedded Device. Multiple pads are present in order to have add-on connection to the sensors, after the electrical components are soldered on |
| Design of the casing that will hold the device. Notice the fillet that allows for the upper part removal and the hole inside the upper cap that allows for wireless signal comunication |
| C:\Users\adrian\AppData\Local\Microsoft\Windows\INetCache\Content.Word\IMG_20170116_152947.jpg  Relative size of the embedded device |

|  |
| --- |
|  |

|  |
| --- |
|  |

## Sensors and the feedback loop

We would like to mention that the decision to include sensors in the design of the device can actually help us increase the duration of the battery. We currently estimate that the most energy expensive procedures will be the electrical stimulation and the Bluetooth communication. While the Bluetooth low energy stack allows us to reduce the power usage, we needed a way to detect objectively a successful stimulation. The PH and humidity sensors provide this mechanism, which, when used in a feedback loop with the neuro-electro-stimulator will provide just the amount of stimulation required, thus reducing battery drain, and improving its life.

# Next Iteration / Future work

After testing out the original version we realized that there is still room for improvement. The most important thing we must improve on is the size, currently the design fits inside a cylindrical case of 13.5x6 mm.

There are 2 designs we are considering at the moment: one with a high degree of success sized 11.5x6 mm, and one that is incomplete, sized ~10x5.5.

Each of these designs come with the assumption that the new design will not have the central hole present, and that a different set of batteries will be used (CR927 – 9x2.7 mm or a set of SR67 – 7.9x1.65, for a total of 7.9x3.3 mm used only by the batteries). The matter of the electronic design is still being internally debated.

These new batteries will require additional phases of testing to see if they actually can provide power for enough time to be viable.

Also, the Android reference application and the firmware must be updated to include the features of

## Sending sensor data in bulk

Currently the Microsal device must be polled in order to obtain the sensor information. This modification would provide great advantages in battery saving, as the device can spend more time in the sleep phase, instead of constantly transmitting sensor data. This improvement would also mean setting up buffers and improving the communication protocol between the Microsal and the Android device.

## synchronization with timestamp

This modification proposes creating a BLE characteristic that would represent an encoded structure that would report the pH and the humidity correlated with timestamp information. We are debating if there would be a need for the device to know the actual time or just the time since boot would be enough. Each approach has limitations and challenges.

## Better battery prediction

Currently the batteries have very flat voltage drop off curves, meaning that without a strong correlation between battery characteristics and measurement capability, we are put in the position where the battery can simply fizzle out with little warning.

## UI improvements and separation of patient/doctor interfaces

The final version of the Android application must have two different UIs, one for the patient, which only reports data for the user’s devices, and does not allow the user to reconfigure his device, and the medical technician version, that has discovery mechanisms and is capable of more advanced probing and configuration of current devices.

## Newer iterations of the design require far smaller dimensions for the device casing, which also means less available space for the batteries and electronics.

The most important constraint and metric for the success of the project remains the size of the device, it must be reduced as much as possible in order to ensure user adoption. Currently, the device is large enough that teeth from lower and the upper jaw need to be sculpted in order for the device to be attached.

## Power optimizations

For example, determine the lowest power configuration that still preserves the neuro-electro-stimulation effect. Another optimization that must be implemented is the monitoring and incorporating feedback loops from the sensors, a reasonable salivary response should not be over-stimulated. The second most important metric for the project’s success, after the device’s size, is the lifespan of the device’s battery. Naturally, a device with a longer span is to be desired, but a minimum battery life of several weeks is required in order to ensure user adoption. Note that the device cannot be recharged, as the electronic logic for charging simply cannot be integrated into such little space.

References

1. *Workshop București 28.10.2016 Abordarea interdisciplinara a pacientului cu xerostomie; perspective actuale de diagnostic si tratament*, *PIM* ed., Iasi, Romania, 2016. ISBN 978-606-13-3276-2*.*
2. Adrian Nitu, Dan Tudose, Dan Dragomir. (2016, Aug). Dental Implant With Contact Nano-Sensors for the Treatment of Xerostomia. Presented at RoEduNet
3. E. E. Harold Warner, E. David Martin, V. Speck, B. Nathan,  “Electrostimulation of erection and ejaculation and collection of semen in spinal cord injured humans”, http://www.rehab.research.va.gov/ jour/86/23/3/pdf/warner.pdf, July 1986.
4. Porter SR, Konttinen YT. Fedele S1, Wolff A, Strietzel F, López RM. Neuroelectrostimulation in treatment of hyposalivation and xerostomia in sjögren’s syndrome: a salivary pacemaker. http://www.ncbi.nlm.nih.gov/pubmed/18671323, August 2008.